# EMSCOPE WEBSOCKET PROTOCOL DOCUMENTATION
This page documents the WebSocket protocol for EMScope software ≥ 4.1.
For software version 4.0, the protocol is identical except that measurement data was always sent as JSON (see section 7).
The Emscope software consists in a WebSocket server listening on port 8010 of the machine and a Web Server listening on the standard HTTP port 80 that serves the front end html and related files.
To configure, parametrize the measurements and manage the standards, a JSON style command set is used through a WebSocket connection.
## 1. First connection
Upon connection, the server does not send any data automatically. The client must first send a session_UUID message to activate itself. See section 8 for details on session locking.
The server responds to this with device info:
{
  "SN": "123456789",
  "MAC": "00:11:22:33:44:55",
  "SFP_SN": "SFP123456",
  "sw_version": "4.1",
  "measurement_uncertainty": "0.5 dB",
  "num_points": 8192
}
After that, the client must send trace_type to enable measurement data streaming. See section 6 for a full initial configuration example.
## 2. Default values
The default values for measurements are:
- Channel: LG
- Detector: PK
- Tracer: Clear/write
- Units: dbuv
- RWB: 9
- Subranges: 10
- Margin: 10
- Active: true
- Mode: circuit
- input_attenuator: auto
- reference_level: 100 (dbuV)
- average: 10
## 3. Commands and responses
The following parameters can be sent to configure the device:
Name	Parameter	Type	Default
Channel	measure_channel	“lg”, “ng”, “cm”, “dm” (RX2) / “l1”, “l2”, “l3”, “n” (RX4)	“lg”
Detector type	detector_type	“pk”, “qp”, “av”	“pk”
Trace	trace_type	“clearwrite”, “maxhold”, “minhold”, “freeze”, “average”	“clearwrite”
RBW	rbw	“200”,“9”,“120” (CISPR), “1”, “10” (MIL), “200_9”, “1_10” (dual band)	“9”
Average (for trace average)	average	integer (10 - 20)	10
Mode	measure_mode	“circuit”, “modal”	“circuit”
Reference level	reference_level	integer	100
Input attenuator	input_attenuator	{0..78} (integer), “auto”	“auto”
Amplitude units	amp_units	“dbm”, “dbmv”, “dbuv”, “watts”, “volts”	“dbuv”
Serial number	SN	string
Sweep time	sweep_time	{1..15} (float)	1
External loss attenuator	external_loss	External loss name (string)	none

## 4. Changing RBW
After setting the RBW we shouldn't send any commands until the device it sends back the RBW value. This is so because a hot-change of firmware, that typically takes around 3.5 seconds, is required. To set a new RBW, send a rbw field with the string corresponding to the RBW:
{
  "rbw": "9", "threephase": false
}
RBW table:
Value	RBW	Frequency	Standard
200	200 Hz	9 kHz - 150 kHz	CISPR 16-1-1
9	9 kHz	150 kHz - 30 MHz	CISPR 16-1-1
120	120 kHz	30 MHz - 110 MHz	CISPR 16-1-1
1	1 kHz	10 kHz - 150 kHz	MIL-STD-461
10	10 kHz	150 kHz - 30 MHz	MIL-STD-461
200_9	200 Hz / 9 kHz	9 kHz - 30 MHz	Dual-band
1_10	1 kHz / 10 kHz	10 kHz - 30 MHz	Dual-band

threephase (boolean, default false) — set to true for three-phase (RX4) instruments and EMI_3P. Must be sent together with rbw in the same message. Sending rbw without threephase defaults to false.
## 5. Additional configuration fields
The following fields can be sent as individual JSON messages at any time after the RBW has been set.
### display_range
Restricts the frequency window of the data sent by the server to a sub-range of the active band.
Send:
{ "display_range": [from_hz, to_hz] }
Both values are absolute frequencies in Hz and must lie within the band of the currently active RBW.
Example — receive only the 150 kHz – 30 MHz portion of the 9 kHz band:
{ "display_range": [150000, 30000000] }
To restore the full band, send the full frequency limits of the active band.
### visible
Controls whether the trace is shown on the instrument's front panel display. Does not affect data transmission to WebSocket clients.
Send:
{ "visible": true }
{ "visible": false }
### freeze (trace_type)
In addition to the trace types listed in section 3, the value “freeze” is also accepted:
{ "trace_type": "freeze" }
The instrument continues sweeping internally but stops updating the displayed trace. Sending a different trace_type unfreezes it.
## 6. Initial configuration
A typical initial configuration parameters example would be:
{
  "detector_type": "pk",
  "measure_channel": "lg",
  "trace_type": "clearwrite",
  "amp_units": "dbmv",
  "rbw": "9",
  "reference_level": 70,
  "input_attenuator": "auto",
  "sweep_time": "1"
}
## 7 - Receiving measurements
To start receiving measurements, send at least trace_type. From that point the server streams data continuously after each sweep completes.
Since SW 4.1, measurement data uses a Binary WebSocket frame.
The frequency axis is sent as a separate JSON message and is cached — it is not repeated in every sweep frame.
### 7.1 - Frequency axis (JSON)
Sent by the server once when the RBW changes or the display range is updated. All subsequent sweep frames share this cached axis until the next such change.
{ "freq": [9000.0, 18000.0, 27000.0, ...], "rbw": "9" }
- freq — array of absolute frequencies in Hz, one per measurement bin.
- rbw — the active RBW string, echoed for reference.
### 7.2 - Binary amplitude frame (EMI mode, SW ≥ 4.1)
One binary WebSocket frame is sent per sweep. The layout is:
Offset Size Field
------ ------- ----------------------------------------
0 1 byte flags
1 3 bytes padding (ignored)
4 N×4 B float32 amplitude values (little-endian)
flags byte:
- flags & 0x01 → overload: the ADC saturated during this sweep. Reduce reference_level or increase input_attenuator.
- flags & 0x02 → repetition: a quasi-peak repeat sweep was triggered (200 Hz and 9 kHz bands only).
The number of float32 values (N) matches the length of the last received freq array. Each value corresponds to the amplitude at freq[i], in the units configured via amp_units.
### 7.4 - Input attenuator echo
When input_attenuator is set to “auto”, the server sends the resolved hardware attenuation as a separate JSON message after each configuration change:
{ "input_attenuator": 30 }
This message is independent of the sweep data frame. It is not embedded in binary frames.
## 7. Locking the device to the current session
Sending session_UUID is required — the server will not activate the client or send any data until it receives this message. It also serves as a session lock: if another client already holds the lock with a different UUID, the server closes the connection with close code 4003. Use the same UUID for subsequent reconnections to the same session.
{
  "session_UUID": "12345qwerty"
}
The server responds immediately with device info (SN, MAC, SFP_SN). Once this exchange is complete, the client is active and can send configuration messages.
## 8. Standards (or Limits)
### 8.1 Getting the standards list
Send the field get_standards with a value of “true”:
{
  "get_standards": true
}
It will return a standards object with an array of standards with the following format:
"Standard name": [ { "rbw": value, "data":
  [ From frequency (Mhz), To frequency (Mhz), From Quasi-peak (dBuV), To Quasi-peak (dBuV), Average from Quasi-peak (dBuV), Average to Quasi-peak (dBuV) ],
  ...
  }
]
For example:
{
  "standards": [
    {
      "Standard One": { "rbw": "9", "data": [
        [1, 2, 3, 4, 5, 6],
        [7, 8, 9, 10, 11, 12],
        ...
      ]
      }
    },
    {
      "Standard Two": { "rbw": "200", "data": [
        [21, 22, 23, 24, 25, 26],
        [27, 28, 29, 210, 211, 212],
        ...
      ]
      }
    },
    ...
  ]
}
### 8.2 Creating a standard
To create a standard, send a name field with the name of the new standard, a modify field set to “false” and a values array in the same format we described before when receiving the standards list.
{
  "name":"New standard",
  "modify":false,
  "standard_rbw": "10",
  "values":[
    ["1","2","3","4","5","6"],
    ["11","12","13","14","15","16"],
    ...
  ]
}
### 8.3 Editing a standard
Send the same values as creating it, but with the field modify set to true and an additional field, original_name, just in case we wanted to rename the original name, i.e:
{
  "original_name": "New standard",
  "name":"Modified standard",
  "modify":true,
  "standard_rbw": "1",
  "values":[
    ["21","22","23","24","25","26"],
    ["11","12","13","14","15","16"],
    ...
  ]
}
### 8.4 Deleting a standard
Simply send a delete_standard with the name of the standard, i.e:
{
  "delete_standard": "Standard One"
}
### 8.5 Resetting standards to factory defaults
Simply send a reset_standards with a value of “true”:
{
  "reset_standards": true
}
## 9. Sessions (or Presets)
### 9.1 Getting the sessions list
Send the field get_sessions with a value of “true”:
{
  "get_sessions": true
}
It will return a sessions object with an array of sessions with the following format:
[
"Session name": [BLOB DATA],
  ...
]
For example:
{
  "sessions": [
    {
      "Session One": "{\"PROGRAM_VERSION\":1.25,\"id_counter\":2,\"current_center_freq_units\":\"mhz\",\"current_start_freq_units\":\"mhz\",\"current_stop_freq_units\":\"mhz\",\"current_span_units\":\"mhz\",\"current_center_frequency\":15000000,\"current_start_frequency\":0,\"current_stop_frequency\":30000000,\"current_span\":30000000,\"measure_mode\":\"circuit\",\"current_display_scale\":0,\"aDetectors\":[{\"type\":\"normal\",\"measure_channel\":\"lg\",\"trace_type\":\"clearwrite\",\"detector_type\":\"pk\",\"amp_units\":\"dbuv\",\"color\":\"red\",\"hidden\":false,\"standard\":null,\"id\":0,\"sweep_time\":1},{\"type\":\"normal\",\"measure_channel\":\"lg\",\"trace_type\":\"clearwrite\",\"detector_type\":\"pk\",\"amp_units\":\"dbuv\",\"color\":\"red\",\"hidden\":false,\"standard\":null,\"id\":1}],\"reference_level\":100,\"reference_level_volts\":0.1,\"amplitude_step\":10,\"aCurrentFreq\":[150000,30000000],\"aCurrentSpan\":[0,100],\"aMarkers\":[[],[]],\"aMMarkers\":[],\"aStandards\":{\"CISPR 22 CLASS A\":[[\"0.15\",\"0.5\",\"79\",\"79\",\"66\",\"66\"],[\"0.5\",\"30\",\"73\",\"73\",\"60\",\"60\"]],\"CISPR 22 CLASS B\":[[\"0.15\",\"0.5\",\"66\",\"56\",\"56\",\"46\"],[\"0.5\",\"5\",\"56\",\"56\",\"46\",\"46\"],[\"5\",\"30\",\"60\",\"60\",\"50\",\"50\"]]},\"aStandardsQuasiPeak\":{},\"aStandardsAverage\":{},\"active_tab\":1,\"ZERO_SPAN_MODE\":false,\"amp_units\":\"dbuv\",\"rbw\":\"9\",\"input_attenuator\":\"auto\",\"sweep_time\":1,\"current_color_theme\":\"dark\"}",
        ...
      ]
    },
    {
      "Session Two": "......",
        ...
      ]
    },
    ...
  ]
}
### 9.2 Creating a session
To create a session, send a session_name field with the name of the new session, a modify field set to “false” and a session_values array in the same format we described before when receiving the sessions list.
{
  "session_name":"New session",
  "modify":false,
  "session_values": ".....",
    ...
  ]
}
### 9.3 Editing a session
Send the same values as creating it, but with the field modify set to true and an additional field, original_session, just in case we wanted to rename the original name, i.e:
{
  "original_session": "New session",
  "session_name":"Modified session",
  "modify":true,
  "values": "......",
    ...
  ]
}
### 9.4 Deleting a session
Simply send a delete_session with the name of the session, i.e:
{
  "delete_session":"Session One"
}
### 9.5 Deleting all sessions
Simply send a reset_sessions with a value of “true”:
{
  "reset_sessions": true
}
## 10. External loss attenuators
### 10.1 Getting the external loss attenuators list
Send the field get_external_losses with a value of “true”:
{
  "get_external_losses": true
}
It will return a external_losses object with an array of external loss attenuators with the following format:
"External loss attenuator name": [ [ Frequency (Mhz), Line (dB), Neutral (dB) ], [ Frequency (Mhz), Line (dB), Neutral (dB) ], ... ]
For example:
{
  "external_losses": [
    {
        "External loss attenuator 1": [ [ 1, 2, 3 ], [ 2, 3, 4 ], ... ]
    },
    {
        "External loss attenuator 2": [ [ 10, 20, 30 ], [ 20, 30, 40 ], ... ]
    },
    ...
  ]
}
### 10.2 Creating an external loss attenuator
To create an external loss attenuator, send an external_loss_name field with the name of the new external loss attenuator, a modify field set to “false” and a values array in the same format we described before when receiving the external loss attenuators list.
{
  "external_loss_name":"New External loss attenuator",
  "modify":false,
  "values":[
    ["1","2","3"],
    ["10","20","30"],
    ...
  ]
}
### 10.3 Editing an external loss attenuator
Send the same values as creating it, but with the field modify set to true and an additional field, original_external_loss_name, just in case we wanted to rename the original external_loss_name, i.e:
{
  "original_external_loss_name": "New External loss attenuator",
  "external_loss_name":"Modified External loss attenuator",
  "modify":true,
  "values":[
    ["11","22","33"],
    ["21","22","23"],
    ...
  ]
}
### 10.4 Deleting an external loss attenuator
Simply send a delete_external_loss with the name of the external loss attenuator, i.e:
{
  "delete_external_loss": "External loss attenuator 1"
}
## 11. Report lists
To generate a report list, we must send 3 values: standard (the standard name string), subranges (number of subranges, an integer) and margin (the margin, an integer), i.e:
{
  "standard": "One standard",
  "subranges": 10,
  "margin": 10
}
The answer to the request will be a boolean frase and a report object with an array of arrays in the following order:
[
Marker (integer),
Frequency in Mhz (float),
Peak level in dBuV (float),
Quasi-Peak level in dBuV (float),
Quasi-Peak limit in dBuV (float),
Distance to Quasi-Peak limit [dB] (float),
Average level in dBuv (float),
Average limit in dBuV (float),
Distance to Average limit [dB] (float),
Channel (string),
Compliance (PASS/FAIL)
]
The frase value is true if at least one of the reported emissions is within margin dB of its applicable limit (i.e. the result is close to the pass/fail boundary). false otherwise.
For example:
{
  "report": [
    [ 1, 5, 116.964216, 116.418941, 119.000000, 10.000000, 116.480995, 109.000000, 10.000000, "L", "PASS" ],
    [ 2, 6, 313.963336, 133.418211, 1123440000, 111220000, 334.440995, 239.020000, 30.000000, "N", "FAIL" ],
    ...
  ],
  "frase": true
}
## 12. Keepalive
The server sends periodic keepalive messages to detect disconnected clients:
{ "ping": true }
The client MUST reply immediately with:
{ "pong": true }
If the client does not respond within the timeout, the server will close the connection. Clients that use the official Python API (EmscopeConnection) have this handled automatically.
## 13. Miscellaneous
### 13.1 Getting the activated licenses
Send the field get_licenses with a value of “true”:
{
  "get_licenses": true
}
As a response you will get a licenses object with an array with the name of each activated licenses:
{"licenses": [ "emi", "osc" ]}
### 13.2 Getting the current PCB and FPGA temperature
Send the field get_temps with a value of “true”:
{
  "get_temps": true
}
As a response you will get a temperatures object with an array with the celsius degrees. The first value is for the PCB, the second one for the FPGA:
{"temperatures": [ 45.12345, 50.12345 ]}